## Sheet2
Prüfung Pipetten
Firma/ Labor
Monat/Jahr
Prüfer
Prüfgrund
Liquid Handling Gerät
Hersteller | Hamiton
Typ
Seriennummer
Liquid Klasse
Kanal | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Nennvolumen | 5
Korrigiertes Volumen
Pipettenspitzen | [8 empty] | Volume [µl] | Tip Size [µl] | Trueness [%] | Precision [%]
Typ/ Hersteller | Hamilton | [7 empty] | 0.5 | 10 | 10 | 6
LOT Nr. | [8 empty] | 1 | 10 | 5 | 4
Maximales Volumen verwendeter Pipettenspitzen [µl] | 50 | [7 empty] | 5 | 10 | 2.5 | 1.5
[9 empty] | 10 | 10 | 1.5 | 1
Messreihe Pipettiervolumen | Gewicht [g] | [7 empty] | 0.5 | 50 | 10 | 6
1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 50 | 5 | 4
2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 50 | 2.5 | 1.5
3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 50 | 50 | 2 | 0.75
4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 300 | 5 | 2
5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 50 | 300 | 2 | 0.75
6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 200 | 300 | 1 | 0.75
7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 1000 | 7.5 | 3.5
8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 100 | 1000 | 2 | 0.75
9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1000 | 1000 | 1 | 0.75
10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Mittelwert [g] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Standardabweichung [g] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Abweichung vom Sollwert [g] | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
ermittelte Richtigkeit [%] | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
ermittelte Präzision [%] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Richtigkeit nach Hersteller [%] | 2.5 | 2.5
Präzision nach Hersteller [%]
Info: Volumen, welche nicht in Spezifikationsliste aufgelistet sind werden nach Spezifikation des nächst höheren Volumen kontrolliert
Liquid Handling Workstation |  |  |  |  | Microlab STAR
Umweltkonditionen |  |  |  |  | Laborumfeld
Aspirationsdetails
Basisinformationen
Tip Use: |  |  |  |  | Single
Pipettiereinstellungen
Pipettenspitzen
Liquid Class
Aspiration mode: |  |  |  |  | Aspiration
Consecutive Aspiration
Aspirate all
Aspiration position:
cLLD |  |  |  |  | Sensitivity:
[5 empty] | Submerge depth: |  |  | mm
pLLD |  |  |  |  | Sensitivity:
[5 empty] | Submerge depth: |  |  | mm
Fix height from bottom
Touch off |  |  |  |  | yes/no |  |  | mm
Side Touch |  |  |  |  | yes/no
Liquid following |  |  |  |  | yes/no
Mix |  |  |  |  | yes/no
Mix cycles
Mix position
Mix volume
Dispensionsdetails
Basisinformationen
Initiale Situation der Ziellabware |  |  |  |  | empty/prefilled
Pipettiereinstellungen
Dispensation mode |  |  |  |  | Jet part volume
Jet empty tip
Surface part volume
Surface empty tip
Drain tip in jet mode
From liquid class definition
Blowout tip
Empty tip at fix position
Aspiration position:
cLLD |  |  |  |  | Sensitivity:
[5 empty] | Submerge depth: |  |  | mm
pLLD |  |  |  |  | Sensitivity:
[5 empty] | Submerge depth: |  |  | mm
Fix height from bottom
Touch off |  |  |  |  | yes/no |  |  | mm
Side Touch |  |  |  |  | yes/no
Liquid following |  |  |  |  | yes/no
Mix |  |  |  |  | yes/no
Mix cycles
Mix position
Mix volume

## Sheet3
Tip Size [µl] | Volume [µl] | Trueness [%] | Precision [%] | Pipettengrößen
10 | 0.5 | 10 | 6 | 10
10 | 1 | 5 | 4 | 50
10 | 5 | 2.5 | 1.5 | 300
10 | 10 | 1.5 | 1 | 1000
50 | 0.5 | 10 | 6
50 | 1 | 5 | 4
50 | 5 | 2.5 | 1.5
50 | 50 | 2 | 0.75
300 | 10 | 5 | 2
300 | 50 | 2 | 0.75
300 | 200 | 1 | 0.75
1000 | 10 | 7.5 | 3.5
1000 | 100 | 2 | 0.75
1000 | 1000 | 1 | 0.75